NMR Magnet Assembly

# 1. Overview

This is a step-by-step guide for building a new NMR magnet using the assembly fixtures made for it. The parts are the magnet core, the permanent magnets, thin steel center plates, thick steel side plates, and a set of HDPE strips.

The magnets are strong and will snap together hard if you let them. The point of the fixtures is to lower each magnet slowly and keep everything lined up. Work carefully and keep your hands out of the gap.

# 2. Safety

Magnets are the main hazard. They can crush a finger, chip or crack on impact, and pull in nearby steel. Keep them under control the whole time.

- Seat magnets with the insertion tool and strips, not with your hands between two magnetic faces.

- Never let a magnet drop in free. Lower it gradually using the strips.

- Keep loose steel tools, screws, and other magnets away from the work area.

- Wear safety glasses. Magnets can fracture and throw pieces if they hit each other.

- Keep phones, drives, cards, and watches away. Anyone with a pacemaker or implant should stay clear.

- Have a second person or third person help hold the fixture and hand off parts during the assembly.

# 3. Parts and Tools

Have everything ready before you start.

| Part | Description |
| --- | --- |
| Magnet core | The block with the magnet size hole. Each magnet sits in this pocket. |
| Magnet Assembly Tool | The fixture with the round boss. It pushes each magnet straight down into the pocket. |
| Side frames (2 sizes) | The open frames with side windows. They hold the stack so you can slide the thick side plates in. Two sizes, one for each plate size. |
| HDPE strips | Thin strips cut from HDPE sheet. You stack them under a magnet and pull them out one at a time to bring them down slowly. |
| Magnets | The NMR magnets |
| Center plates | Thin steel discs cut to the magnet size and ground flat. They go against magnet faces to make the field more even. |
| Side plates | Thick steel plates that go on the four sides at the end to keep the field inside. Two sizes. |

# 4. Prepare the Center Plates

Do this before handling any magnets.

- Cut the thin steel plates to the magnet size.

- Grind or lap each plate flat and smooth so the magnet seats fully against it.

- Deburr the edges, wipe the plates clean, and set them aside away from the magnets.

A burr or bit of dirt under a plate leaves a small gap that throws off the field, so it is worth getting these flat.

# 5. Build the Magnet Stack

Work on a flat clean surface with the core pocket facing up. Keep the insertion tool and strips within reach and wear safety glasses.

- Place a center plate flat in the bottom of the pocket. Check it sits flush with no rocking.

- Lay HDPE strips across the pocket to make a platform that holds the magnet above the plate.

- Load the magnet onto the magnet assembly tool so it stays square. Keep your hands on the tool, not in the gap.

- Lower the magnet with the tool until it rests on the strips, not yet on the plate.

- Pull the strips out one at a time. The magnet steps down a little with each one until it sits on the plate.

- Check the magnet is seated flat and centered.

- Place the next center plate on the other side of the core and lower it by using the same magnet assembly tool.

- Lower the second magnet the same way. Check its orientation before bringing it down.

- Repeat for any more magnet and plate layers the design needs.

# 6. Add the Side Plates

With the magnets and plates stacked, add the thick steel plates on the four sides using the frames.

- Pick the frame that matches the plate size (there are two).

- Set the stack in the frame so a side window line up with the side you are filling.

- If needed, use strips to hold the plate slightly out so it goes in under control instead of slamming.

- Pull the strips and slide the plate against the stack until it sits flat.

- Do all four sides, switching frame size when the plate size changes.

- Check each plate is seated against the stack with no gaps. Use the dead blow hammer to adjust the plates.

# 7. Final Check

- Confirm every layer is flat and centered and all four side plates are flush. Use dead blow hammer if needed.

- Look for any pinched strip, debris, or gap between faces.

- Measure the field with a gaussmeter and compare it to what you expect for this magnet.